## Failure Components
Failure Date | Time to Failure | Hutch | Serial Number W/O Production Number | Symptoms | Failure Cause/Correlation | Total Inventory | Production Number | Deployment Date
2024-07-04 | 36 | TMO | SLC1760ds | Unresponsive stage AND Piezo damaged | Large spike in MCS2 module temperature (65°C) |  | 89
2024-05-26 | 40 | TMO | SLC1740ds | Unresponsive stage, Hanging after sticky movements AND Piezo damaged | User error of stage being indefinitely run into the EJX enclosure |  | 85
2024-06-28 | 33 | TMO | SLLV42lc | Not "homing" properly | Sensor or encoder strip damage |  | 469
2023-11-27 | 34 | TMO | STT25.4v | Unresponsive tilt axis AND Piezo damage | Power outage |  | 21
2023-02-05 | 39 | TMO | SLLV42sc | Unresponsive linear stage AND Piezo damage | Power outage |  | 463
2023-11-27 | 37 | TMO | STT25.4v | ATM_MR0_TILT1 unresponsive AND Piezo damage | Unknown
2023-11-27 | 37 | TMO | STT25.4v | ATM_MR1_TILT1 unresponsive AND Piezo damage | Unknown
2024-01-30 | 38 | TMO | SLC1780ds | Absolute positions change unexpectedly | Calibration error caused by firmware update
2024-01-30 | 37 | TMO | CLL42lc | Absolute positions change unexpectedly | Calibration error caused by firmware update
2024-02-27 | 35 | chemRIXS | STT50.8iv | Intermittent sensor failure on encoded tilt axis | Excessive weight of optic causing piezo slippage
2024-02-05 | 32 | chemRIXS | STT25.4v | Unresponsive axes | Power outage
2024-02-05 | 38 | chemRIXS | CLL42lc | Stage failing sporadically in closed-loop movement | Sensor-type incompatibility with an L-type sensor on a sensor module with I-type sensors |  | 47
2024-04-22 | 40 | chemRIXS | SLC2445ds | Stage failing sporadically in closed-loop movement AND Not "homing" properly | Wrong positioner type assigned AND Sensor-type conflict with neighboring stages
2024-06-12 | 33 | chemRIXS | SLLV42lc | Unresponsive movement AND Piezo Damage | Unknown |  | 473
2024-02-29 | 37 | chemRIXS | STT50.8iv | Unknown | Unknown |  | 6
2024-02-14 | 39 | chemRIXS | CLL42lc | Unknown | Unknown |  | 47
2024-03-18 | 32 | TMO | STT25.4v | Unknown | Unknown |  | 21
2024-03-18 | 38 | TMO | SLLV42sc | Unknown | Unknown |  | 463
2024-02-14 | 34 | chemRIXS | STT25.4v | Unknown | Unknown |  | 24
2024-01-01 | 34 | TMO | STT25.4v | Unknown | Unknown |  | 27
2023-11-22 | 34 | Investigating | STT25.4v | Piezo burned | Unknown |  | 25
2023-11-22 | 35 | Investigating | STT50.8v | Cable Damaged | Unknown |  | 6
2023-11-22 | 39 | Investigating | STT50.8v | Piezo burned | Unknown |  | 29
2023-11-22 | 30 | Investigating | SLC1780ds | Unknown | Unknown |  | 40
2023-11-22 | 33 | Investigating | SR7012l | Piezo burned | Unknown |  | 18
2024-06-12 | 37 | chemRIXS | SLLV42lc | Unknown | Unknown |  | 473