## Failure Components
Purchase Date | Failure Date | Hutch | Serial Number Without Production Number | Production Number | Symptoms | Failure Cause/Correlation | Total Inventory
| 2024-07-04 | TMO | SLC1760ds | 89 | Unresponsive stage AND Piezo damaged | Large spike in MCS2 module temperature (65°C)
| 2024-05-26 | TMO | SLC1740ds | 85 | Unresponsive stage, Hanging after sticky movements AND Piezo damaged | User error of stage being indefinitely run into the EJX enclosure
| 2024-06-28 | TMO | SLLV42lc | 469 | Not "homing" properly | Sensor or encoder strip damage
| 2023-11-27 | TMO | STT25.4v | 21 | Unresponsive tilt axis AND Piezo damage | Power outage
| 2023-02-05 | TMO | SLLV42sc | 463 | Unresponsive linear stage AND Piezo damage | Power outage
| 2023-11-27 | TMO | STT25.4v |  | ATM_MR0_TILT1 unresponsive AND Piezo damage | Unknown
| 2023-11-27 | TMO | STT25.4v |  | ATM_MR1_TILT1 unresponsive AND Piezo damage | Unknown
| 2024-01-30 | TMO | SLC1780ds |  | Absolute positions change unexpectedly | Calibration error caused by firmware update
| 2024-01-30 | TMO | CLL42lc |  | Absolute positions change unexpectedly | Calibration error caused by firmware update
| 2024-02-27 | chemRIXS | STT50.8iv |  | Intermittent sensor failure on encoded tilt axis | Excessive weight of optic causing piezo slippage
| 2024-02-05 | chemRIXS | STT25.4v |  | Unresponsive axes | Power outage
| 2024-02-05 | chemRIXS | CLL42lc | 47 | Stage failing sporadically in closed-loop movement | Sensor-type incompatibility with an L-type sensor on a sensor module with I-type sensors
| 2024-04-22 | chemRIXS | SLC2445ds |  | Stage failing sporadically in closed-loop movement AND Not "homing" properly | Wrong positioner type assigned AND Sensor-type conflict with neighboring stages
| 2024-06-12 | chemRIXS | SLLV42lc | 473 | Unresponsive movement AND Piezo Damage | Unknown
| 2024-02-29 | chemRIXS | STT50.8iv | 6 | Unknown | Unknown
| 2024-02-14 | chemRIXS | CLL42lc | 47 | Unknown | Unknown
| 2024-03-18 | TMO | STT25.4v | 21 | Unknown | Unknown
| 2024-03-18 | TMO | SLLV42sc | 463 | Unknown | Unknown
| 2024-02-14 | chemRIXS | STT25.4v | 24 | Unknown | Unknown
| 2024-01-01 | TMO | STT25.4v | 27 | Unknown | Unknown
| 2023-11-22 | Investigating | STT25.4v | 25 | Piezo burned | Unknown
| 2023-11-22 | Investigating | STT50.8v | 6 | Cable Damaged | Unknown
| 2023-11-22 | Investigating | STT50.8v | 29 | Piezo burned | Unknown
| 2023-11-22 | Investigating | SLC1780ds | 40 | Unknown | Unknown
| 2023-11-22 | Investigating | SR7012l | 18 | Piezo burned | Unknown
| 2024-06-12 | chemRIXS | SLLV42lc | 473 | Unknown | Unknown